# Private Testing Desktop Review (stage 2) Operating Process for SARS-CoV-2 Tests
Published 28th July 2021
This guidance outlines how the government will undertake the desktop review of COVID-19 viral detection tests submitted for assessment.

## Overview of the Process
This review is a systematic assessment of the evidence submitted by a supplier against a minimum required data set. The purpose of this step is to prevent products that are below the expected standards from progressing to a technical validation in a laboratory. The assessment will be undertaken by a suitably qualified scientific advisor against the information submitted. The main areas of assessment are as follows:
1. Manufacturer and product information
2. Regulatory status
3. Intended Use Case
4. Product performance
5. Biosafety

### Step 1: Manufacturer and product information
As a minimum the applicant should supply a Word or PDF document of the Instructions for Use (IFU) for the product. This document will be required to demonstrate the product aligns with the intended use and evidence the stated performance characteristics of the product. Additional supporting publications will be most likely required to evidence the full characteristics of the submitted product

The submission is assessed to ensure that the key documents for the application are present.

Supplementary documents uploaded with application	Assessment	Response
Current version of IFU	Supplied/Not Supplied	Automatic reject if not supplied
IFU for comparator assay that was used to calculate performance characteristics	Supplied/Not Supplied	Automatic reject if not supplied
Biosafety documents (including viral inactivation where a claim of inactivation is stated in the IFU)	Not applicable/ / Supplied/Not Supplied	Automatic reject if not supplied
Evidence of performance characteristic including materials and method for limit of detection and Raw CT data per gene target for the comparator RT-PCR results and the Raw CT data for the submitted assay	Supplied/Not Supplied	Automatic reject if not supplied
Field safety notices related to the submitted product issued from 01/01/2020 to date of application	Supplied/Not Supplied	SA to assess impact and resolution of issue
Regulatory certification e.g. CE IVD, UKCA, MHRA derogation	Supplied/Not Supplied	Automatic reject if not supplied

### Step 2: Regulatory status
Regulatory approval	Desirable	Minimum requirement
	UK Certificate of Approval (CA) or Certificate European (CE)	Alternatives – Health Institution Exemptions or exceptional use authorization (i.e. MHRA derogation)

From 1 July 2023, a UKCA mark will be required in order to place a device on the Great Britain market. Prior to this date, manufacturers can use the UKCA mark on a voluntary basis and the CE mark will continue to be accepted on the Great Britain market until 30 June 2023.

You will need to supply a copy of the certificate and any exemption reports

The regulatory approval documentation will be assessed for the following;

Review area	Documentation	Evidence	Guidance Notes
Regulatory status at time of application	Regulatory / Certificate	Automatic reject if not supplied	Product must have either CE IVD, UKCA or MHRA exemption. Products that do not meet this criteria will be automatically rejected

### Step 3: Intended Use Case
The IFU will be assessed for:

Review area	Documentation	Evidence	Guidance Notes
Intended use	Please attach and highlight relevant reference in IFU and publications	Automatic reject if not supplied	Intended use case as described in the IFU
Target population	Please attach and highlight relevant reference in IFU and publications	Automatic reject if not supplied	Please describe if your product in intended for symptomatic and/or asymptomatic cases
Sample collection type and requirements	Please attach and highlight relevant reference in IFU and publications	Automatic reject if not supplied	Please state the specifications and product refence numbers for collection devices (including swabs, media and containers) that have been validated for use with your product
Target use setting	Please attach and highlight relevant reference in IFU and publications	Automatic reject if not supplied	Please state if your product the assay is intended for point of care, near patient testing and/or laboratory based
Target user	Please attach and highlight relevant reference in IFU and publications	Automatic reject if not supplied	Trained healthcare professional member of the allied health profession, trained lay-person self-test,
Please state the time to result for your product	Please attach and highlight relevant reference in IFU and publications	Automatic reject if not supplied	Time to result excludes delivery and receipt of samples but is inclusive of relaying of results into LIMS system

### Step 4: Product performance
#### Minimum Evidence to support performance claim
Desirable	Essential
150 clinical "Positive"; 250 clinical "Negative" as determined by a comparator laboratory RT-PCR	100 clinical "Positive" and150 clinical "Negative" as determined by a comparator laboratory RT-PCR

The comparator method must meet the the diagnostic criteria from the TVG Test Thresholds guidance.[LINK]
Clinical diagnosis or contrived samples are not acceptable comparator methods.
The performance results raw data (CT values or equivalent) is required in the supplementary evidence submitted via the portal to evidence that the dynamic range of the assay has been evaluated.

### Clinical Performance
The IFU, relevant raw data and published supporting documents will be assessed for key requirements for evidence of clinical performance.
The clinical performance characteristics of the assay. Performance characteristics must be based on the comparison of the submitted product with results obtained using an RT-PCR method that meets the ‘Desirable criteria’ of the MHRA Laboratory TPP
Review area	Documentation	Evidence	Guidance Notes
The comparator assay name, product number and IFU version	Please attach and highlight relevant reference in IFU and publications	Automatic reject without IFU	The comparator assay must itself meet the requirements of the laboratory based MHRA TPP
Sample size of clinical positive samples	Please attach and highlight relevant reference in IFU and publications	Automatic reject if <100	Minimal acceptable number is 100, desired number is greater than or equal to 150
Sample size for negative clinical samples)	Please attach and highlight relevant reference in IFU and publications	Automatic reject is <150	Minimal acceptable number is 150, desired number is 250
Viral load or CT values of positive samples obtained in the comparator assay	Please attach and highlight relevant reference in IFU and publications	Automatic reject without raw data	The CT or viral load of the positive clinical samples should span the dynamic range of the assay
Please state the clinical sensitivity of your assay determine by comparison with an RT-PCR assay and provide raw data to evidence the statement at 106 to 105 copies per ml or equivelent to CT≤25	Please attach and highlight relevant reference in IFU and publications	Automatic reject without raw data	Raw data of the comparator and submitted assay is required to assess the validity of claim
Please state the clinical sensitivity of your assay determine by comparison with an RT-PCR assay and provide raw data to evidence the statement across the dynamic range of the assay (High/ Med/Low equivalent to CT <25, 25 to 30, >30-35)
Please state the clinical specificity of your assay and provide raw data to evidence the statement	Please attach and highlight relevant reference in IFU and publications	Automatic reject	Raw data of the comparator and submitted assay is required to assess the validity of claim
What evidence do you have of the performance of your assay against contrived samples?	Raw data and relevant published papers	for information only	not required but desirable as prevalence changes impact on access to positive samples
Please state the number and name of SARS-COV-2 gene targets for the assay?		for information only

### Analytical Performance
The IFU, relevant raw data and published supporting documents will be assessed for key requirements for evidence of analytical performance.
The analytical performance characteristics of the assay. Performance characteristics must be based on the comparison of the submitted product with results obtained using an RT-PCR method that meets the ‘Desirable criteria’ of the MHRA Laboratory TPP
Review area	Documentation	Evidence	Guidance Notes
Please state Limit of Detection of your submitted product in copies per ml. Where IFU states another unit e.g. copies/reaction please translate to copies/ml	Please attach and highlight relevant reference in IFU and publications	Automatic reject if not stated in copies/ml	Please include the materials and method used to determine limit of detection
Please confirm in silico exclusivity of your product (NAAT based tests only)	Please attach and highlight relevant reference in IFU and publications	Automatic rejection if not supplied	Please provide details of in silico tests for primer and probe sequences
For molecular tests the supplier must provide evidence of regular in silico analysis against SARS-COV-2 variants of concern database as published on gov.uk (https://www.gov.uk/government/publications/covid-19-variants-genomically-confirmed-case-numbers/variants-distribution-of-cases-data). (The supplier must inform PTVG within 48 hours if a variant of concern (VOC) affects the assay target region/s by in silico analysis during validation.)	Please attach in silico evidence	Automatic rejection if not supplied if molecular assay	Please provide details of in silico tests for primer and probe sequences
For antigen based assays, please state the protein target gene/s and region/s			Please provide details of your assay design and target region/s e.g. monoclonal, polyclonal terminus and target protein
In vitro or in silico cross reactivity with common respiratory targets	Please attach and highlight relevant reference in IFU and publications	Automatic rejection if not supplied	In Silico or In vitro studies with clinical or contrived samples known to contain other respiratory virus. Examples include commercial EQA and control panels
In vitro effect of common interferants	Please attach and highlight relevant reference in IFU and publications	Automatic rejection if not supplied	Ref to MHRA TPP interfering substances list

### Technology Threshold performance
The Private Testing Validation process requires key mainstream technology groups to meet a minimum performance threshold reflecting standards identified for the technology area identified. [Add link to threshold guidance]

Where an applicant considers this performance requirement cannot be met due to exceptional features of the product that mitigates this requirement supporting evidence will need to be supplied at the time of submission for consideration.

Requirement	PCR & Isothermal (Extracted)	PCR & Isothermal (Direct)	Antigen
Minimum Performance category expectation*	Sensitivity with 95% two-sided confidence interval entirely above 93% / Specificity with 95% two-sided confidence interval entirely above 97%	Sensitivity with 95% two-sided confidence interval entirely above 70% / Specificity with 95% two-sided confidence interval entirely above 93%	Sensitivity with 95% two-sided confidence interval entirely above 60% / Specificity with 95% two-sided confidence interval entirely above 93%

* Setting a standard for technology types so government does not endorse products that are substandard for their type.

Review area	Documentation	Evidence	Guidance Notes
Technology identified meets key performance requirement	Supplier Confirmation or supporting documents	Automatic reject if not stated	Confirmation that product meets requirement or Mitigating document as required
Documented performance against minimum standards for Sensitivity, Specificity	Please attach and highlight relevant reference in IFU and publications	Automatic reject if not stated	Additional evidence if required or SA assessment of IFU

### Step 5: Biosafety
Applicant must state the biosafety requirements of their submission.
Where the submitted assay includes the use of a reagent that claims to inactivate virus, the manufacturer must be able to evidence according to the BSI safety standards and PHE position statement on inactivation (position-statement-regarding-covid-19-assays-evaluated-by-phe) and safe handling (link).
The IFU and any clearly identified supporting documents will be assessed for:
Review area	Documentation	Evidence	Guidance Notes
Please state the biosafety containment requirements for your submission	Please attach and highlight relevant reference in IFU and publications	Supplied/ / Not Supplied	The biosafety of product will impact the use-case scenario in which your product can be utilised
Does your product include a viral inactivation step	Please attach and highlight relevant reference in IFU and publications	Supplied/ / Not Supplied	Viral inactivation statements are not acceptable without documented evidence. Equivalence to similar products are not acceptable / British Standard for quantitative suspension tests for evaluation of virucidal activity, which requires that disinfectant products demonstrate a titre reduction of at least 4 log in virus titre[1]
If yes, have you provided documented evidence. Please provide documented efficacy evidence according to BSI standards	Please attach and highlight relevant reference in IFU and publications	Not mandatory / Supplied/ / Not Supplied	see / position-statement-regarding-covid-19-tests-evaluated-by-phe / https://www.gov.uk/government/publications/wuhan-novel-coronavirus-guidance-for-clinical-diagnostic-laboratories/wuhan-novel-coronavirus-handling-and-processing-of-laboratory-specimens
If your submission does not include inactivation of clinical samples, can your product be used outside of a laboratory setting	Please attach and highlight relevant reference in IFU and publications	Supplied/ / Not Supplied	Assessment of sample preparation and processing method ,e.g. non-aerosol creating, need for full PPE
[1] British Standards Institution, Chemical disinfectants and antiseptics – Quantitative suspension test for the evaluation of virucidal activity in the medical area – Test method and requirements (Phase 2/Step 1). 2019, BSI Standards Limited

© Crown copyright 2019
Insert name of Directorate/Division/Branch
www.gov.uk/dhsc
This publication is licensed under the terms of the Open Government Licence v3.0 except where otherwise stated. To view this licence, visit nationalarchives.gov.uk/doc/open-government-licence/version/3
Where we have identified any third party copyright information you will need to obtain permission from the copyright holders concerned.